Temperature Calibration for GSL1100X Tube furnace

Please be advised that there is temperature difference between inside tube and outside tube. In order to achieve accurate firing temperature form inside tube furnace, you shall calibrate your tube furnace and set up temperature off from controller before use.  (Please click here to see how to calibrate tube furnace ).

Typical temperature calibration curve and temperature profile for GLS1100X tube furnace are listed as the below, which is made in air without sample load. The result shows that temperature difference varies from + 8.2 oC @600 oC, and  0.5 oC @ 800 oC to - 8oC@1100 oC without setting temperature off.  Constant temperature within +/- 1oC  is about 80 mm, which may vary with temperature and sample size.

The tube furnace is shipped without temperature off-set  (e.g. SC value set at 0 ).  It is your responsibility to calibrate furnace at target temperature and place your sample in correct position  MTI is not responsible for any damage caused by misuse

If you have any difficulty to calibrate the tube furnace , please contact  MTI engineering at  info@mtixtl.com

MTI Corporation                                                                      www.mtixtl.com

860 South 19th street , Richmond, CA 94804, USA

Tel: (510)525-3070     Fax: (510) 525-4705         E-mail: info@mticrystal.com